## Slide 1: Spectrophotometer
Dr Sumitra G,
Professor of Biochemistry,
PSG Institute of Medical Sciences and Research, Coimbatore.
STANBIO ONLINE PG CRASH COURSE 2024
11 September

## Slide 2: Outline
Principle
Components
Applications
Advantages & Limitations
Performance characteristics

## Slide 3: Photometry
Photometry is defined as the measurement of light
Spectrophotometry is defined as the measurement of the intensity of light at selected wavelengths

## Slide 4: Principle

## Slide 5: Beer Lambert’s law
T= I
I0

## Slide 6: Absorbance
Absorbance = log10 1/T
              = log10 100/%T
              = 2- log10 (%T)
T% | A
100 | 0
10 | 1
1 | 2
0.1 | 3
0.01 | 4
0.001 | 5

## Slide 7: Concentration and absorbance
A = a b c
A= Absorbance a = Constant – Molar absorptivity b = path length in cm c = Concentration of the solute in mol/L

## Slide 8: Assumptions/ Limitations
Beer’s law is followed only if the following conditions are met:
Incident radiation on the substance of interest is monochromatic.
The solvent absorption is insignificant compared with the solute absorbance.
The solute concentration is within given limits.
An optical interferant is not present.
A chemical reaction does not occur between the molecule of interest and another solute or solvent molecule.

## Slide 9
Why can’t molar absorptivity be used instead of calibration for all assays?

## Slide 10: Instrumentation

## Slide 11: Components

## Slide 12: Types of spectrophotometers
UV-VIS Spectrophotometer
IR Spectrophotometer

## Slide 13: Light source
The goal is to have a light source that provides high-intensity light uniformly across the spectrum of interest
  bright across a wide wavelength range
  stable over time
  long service life
  low cost
Incandescent
Gas discharge
LED
LASER

## Slide 14: Incandescent lamps
A filament is heated to high temperature using electric current – emits light
Continuous spectrum of light
Examples
  Quartz-halogen lamp – Tungsten - Vis and near IR
  Nernst glower - oxides like zirconium oxide (ZrO₂) and yttrium oxide (Y₂O₃) - IR
  Globar Silicon carbide (SiC) - IR
  Nichrome lights -alloy of nickel and chromium - IR
Notes: Emit light due to thermal excitation. Relation between temperature and spectrum

## Slide 15: Quartz halogen lamp

## Slide 16: Gas discharge- Arc lamps
Arc lamps generate light through an electric arc formed between two electrodes in a gas or vapor.
The high temperatures in the arc cause the gas to emit light.
continuous spectrum – UV, Vis and near IR
Xenon arc lamps, Deuterium arc lamps

## Slide 17
Deuterium lamp
Xenon arc lamp

## Slide 18: Gas discharge- vapour lamps
Low pressure mercury vapour lamp
When an electric current passes through the gas, it excites the atoms, causing them to emit light.

## Slide 19: Light Emitting Diodes

## Slide 20: LEDs

## Slide 21: LEDs

## Slide 22: LASERS
Monochromatic light
Narrow beam width
Versatile pulse widths
Continuous wave dye laser – Argon laser with fluorescent dyes

## Slide 23: Wavelength selector
Required for isolating desired wavelengths – Fixed or scanning.
Filters – Interference filters, absorption filters
Dispersive elements – Prisms, Diffraction gratings
Purity of the exiting light is described as spectral bandwidth

## Slide 24: Absorption filter
Absorption filters work by absorbing specific wavelengths of light and allowing the rest to pass through. Made of colored glass or plastic materials that absorb certain parts of the spectrum.
Advantages:
Durable and easy to use.
Cost-effective for broad applications.
Disadvantages:
Lower spectral resolution.
Not suitable for precise wavelength selection.

## Slide 25: Interference filter

## Slide 26: Interference filter
Use the phenomenon of constructive and destructive interference to transmit light within a narrow wavelength range. Made by stacking thin layers of dielectric materials and metal coatings. Provide high spectral resolution with very narrow bandwidth.
Advantages:
High precision and narrow bandwidth.
Ideal for applications requiring specific wavelength isolation, such as fluorescence spectroscopy and monochromatic light.
Disadvantages:
More expensive than absorption filters.
Low peak transmittance value

## Slide 27: Prism
Principle:
  Prisms disperse light by refraction; different wavelengths bend at different angles due to varying refractive indices.
  Commonly made from glass, quartz, or fluorite (for visible and UV light), and calcite (for infrared).
Advantages:
  Durable and suitable for broad-spectrum use.
Disadvantages:
  Non-linear dispersion; shorter wavelengths (UV) are refracted more than longer wavelengths (IR).

## Slide 28: Diffraction Gratings

## Slide 29: Diffraction Gratings
  Disperse light through interference; light waves are diffracted at different angles based on the wavelength.
  Composed of closely spaced grooves or lines.
Advantages:
  High resolution and precise wavelength selection.
  Suitable for UV, visible, and IR spectral ranges.
Disadvantages:
  Higher-order diffraction can cause overlapping and needs filtering.

## Slide 30: Cuvettes
Round
Square
Rectangle
Glass
Plastic
Quartz

## Slide 31: Photodetectors
Goal of photodetector
Responsive to low radiant energy
Cover broad wavelength
High signal to noise ratio

## Slide 32: Photomultiplier tube

## Slide 34: PMT
Advantages:
  High Sensitivity
  Fast Response Time
  Wide Spectral Range
  Long operational lifespan
Disadvantage
  Very sensitive to stray light

## Slide 35: Photodiode

## Slide 36: Photodiode array
Advantages
Parallel Detection
Sensitivity
Fast Response Time
Disadvantage
Not suited for measuring when light intensity is low

## Slide 37: Charge-Coupled Device
Consists of an array of light-sensitive elements (pixels) that transfer charge to an output device
High Sensitivity
High Resolution
Uniform Response
Disadvantage: Readout Speed low

## Slide 38: Process flow in CCD
Light to charge conversion
Charge accumulation – potential well
Charge transfer
Charge to voltage conversion
Amplification

## Slide 40: CMOS
Photon Detection: Light is converted into charge.
Charge Accumulation: Electrons are collected and stored.
Signal Amplification: The charge signal is amplified within each pixel.
Charge-to-Voltage Conversion: Charge is converted into a voltage signal.
Analog-to-Digital Conversion: The voltage signal is digitized.
Readout: Digital signals from each pixel are read out.
More compact, faster readout speed, low cost but higher noise level compared to CCDs

## Slide 45: Applications

## Slide 46: Applications
Identifying a substance
Quantifying a Molecule
Study enzyme kinetics
Study confirmational changes
Measure purity of proteins and Nucleic acid
Nucleic acid and proteins melting temperature
Colour measurement

## Slide 47
Clinical laboratory – measurement of analytes, drugs concentrations in biological fluids
Molecular biology – Study the quantity, purity and structural changes in DNA
Pharmaceutical industry – verify purity, analyse concentration and quality control
Environmental testing – water and soil quality
Food and beverage industry – food content analysis, colour measurement
Forensic – Identification of substances
Spectrophotometer is an integral part of other instruments like HPLC, ELISA readers, ABG systems for co-oximetry

## Slide 48: Advantages and Limitations

## Slide 49: Advantages
High Sensitivity: Capable of detecting very low concentrations of analytes.
Non-Destructive: Does not alter or destroy the sample during analysis.
Wide Applicability: Can be used for a variety of compounds in different industries (e.g., chemistry, biology, environmental science).
Quantitative and Qualitative: Allows both the identification and quantification of substances.
Fast and Simple: Provides quick results with straightforward procedures.
Automation Capability: Easily integrated into automated systems for high-throughput analysis (e.g., in clinical labs, HPLC detectors, or ELISA readers).

## Slide 50: Limitations
Stray Light: Causes errors at high absorbance values, leading to inaccurate readings.
Sample Preparation: Requires careful preparation; poor handling can affect accuracy.
Turbidity: Suspended particles can scatter light, causing measurement errors.
Beer's Law Deviations: Nonlinear results at high concentrations due to molecular interactions.
Limited Wavelength Range: A single spectrophotometer may not cover all required wavelengths (UV, visible, or IR).
Non-Specific Detection: Overlapping absorbance spectra can make it difficult to differentiate components in mixtures.
Environmental Sensitivity: Temperature, light, and vibrations can introduce noise or variability.

## Slide 51: Performance characteristics

## Slide 52: Performance characteristics
Wavelength accuracy
Spectral bandwidth
Straylight
Linearity
Photometric accuracy
Testing frequency- no regulations regarding - testing frequency monthly testing is preferred and frequency can be increased in high sample throughput environment

## Slide 53: Wavelength accuracy
The wavelength accuracy test is used to ensure that the wavelength axis of the UV-Vis spectrum is accurate (correct and within acceptable limits) across the intended operational range.
This is assessed by using either a sample or a glass filter containing a series of very sharp peaks or by measuring the emission from a lamp.
Deuterium lamp
Mercury emission source
Glass and liquid filters such as holmium oxide, didymium

## Slide 54: Wavelength accuracy using lamp emissions
Deuterium lamp emission: 656.1 nm emission line

## Slide 55: Wavelength accuracy using glass and liquid filters
Notes: Samarium perchlorate UV-VIS and Rare earth sulphate UV

## Slide 56: Photometric Accuracy
These tests are used to ensure that a UV-Vis spectrophotometer yields reliable quantitative measurements.
Neutral density glass filters
Potassium dichromate solution
Nicotinic acid solution

## Slide 57: Photometric linearity
Photometric linearity affects how accurately an instrument measures absorbance with increasing optical density or concentration.

## Slide 58: Stray light
Stray light can be defined as the amount of unwanted wavelengths that are present in the sample beam.
Scattering or diffraction by monochromator, other components of spectrophotometer or sample itself
Light leaks
Fluorescence of the sample
The best achievable stray light value will be around 0.01%T

## Slide 59: Stray light measurement
50 g/L sodium nitrite for 340 nm and 370 nm
10 g/L sodium iodide for 220 nm

## Slide 60: Spectral resolution
The resolution of a UV-Vis spectrophotometer is the narrowest spectral bandwidth that the instrument can achieve.
Toluene in Hexane solution

## Slide 61: Other performance characteristics
Noise: Noise refers to random fluctuations in the signal that are inherent in any measurement system. It is usually a combination of electrical noise from the detector, thermal noise, and fluctuations in the light source.
Baseline flatness
Absorbance stability (drift)
Reproducibility

## Slide 62: Questions?